## Risk analysis
| Title: |  | Risk Management of [Design OR Use and Application OR Production Process; select as appropiate]
| Document No: |  | FRM4503 FMEA Sheet
| Version: |  | 1
| Effective date: |  | 18-Feb-2020
|  | Function: | Name: | Signature: | Date:
| Author: | Software Manager | Faris Naji
| Approval by: | Head DAS | Rik de Wijn
| Subject of Risk Analysis: |  | software module dasrunqc
Include a medical specialist at least for initial and final version of application and design FMEAs of products.
| Risk Management Team: |  | Rik de Wijn
Faris Naji
Risk Management:
For each topic we have listed what could go wrong ("Failure Modes"), the effects this could have on the safety and effectiveness of the Product as well as on patient, what has been done to
reduce the Risk, and what still needs to be done to continue reducing the Risk. Subsequently, we have estimated the likelihood of the Failure, its Detectability in advance (before the Failure
occurs) and its maximum effect on the Safety and Effectiveness of the Product or on the Patient.
Scores used:
P = Probability of Occurrence, the likelihood the mentioned effect (Failure) will occur
D = Detectability, the likelihood the failure will be detected
S = Severity, the effect of the mentioned Failure on the Safety and Effectiveness of the Product patient
RPN = Risk Priority Number (P x D x S)
Maximum value for the Risk Priority Number: 50
Maximum acceptable value for the Risk Priority Number: 10
Number of components with a Risk Priority Number above the permitted limit: [Fill in here]
Use this FMEA sheet in combination with SOP4500 - Risk Management Process and POL4500 - Risk Acceprability Policy
For additional Risk Management cycles, add additional columns as needed.
See sheet "Explanation Scoring System" for scoring scales and Risk Acceptabiliy Matrix.
| Changes: | Changes to the Design FMEA - see sheet "Change History"
Row nr | Part or Process | Topic | Failure mode | Mechanism and cause of failure | Effect of the failure on the function of the product | Effect of the failure on the patient | PDS | RPN (PxDxS) | Risk Control | Actions performed | PDS | RPN (PxDxS) | Risk Control | Actions performed | Remarks | Residual Risk & Acceptability
2 | 2 |  | none
3 | 03A |  | programming error | code | incorrect results calculated | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | input/putput testing | unit test | P:1 D:1 S:5 | 5
4 | 5 |  | none |  |  |  |  | 10
5 | 11A and 11B |  | none
6 | 014A and 014B
7 | 017A1 | data | output has incorrect format | code | incorrect results calculated | worst case: wrong diagnosis | P:1 D:2 S:5 | 10 | unit testing | add a unit test | P:1 D:1 S:5 | 5
8 | 018A11, 018A12, 018A13,018A16 |  | programming error | code | incorrect results calculated | wrong diagnosis | P:2 D:2 S:5 | 20 | unit testing | add a unit test | P:1 D:1 S:5 | 5
9 | REQ011A | concordance result | design is not suitable for cioncordance calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
10 | REQ011B | CV result | design is not suitable for CV calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
11 | REQ014A1 | design of simple table | no risk |  |  |  | P: D: S: |  |  |  | 1
12 | REQ014A2 | design of simple trable | no risk |  |  |  | P: D: S: |  |  |  | 5
13 | REQ014A3 | design of simple trable | no risk |  |  |  | P: D: S: |  |  |  | P: D: S:
14 | REQ014A4 | design of simple trable | no risk |  |  |  | P: D: S: |  |  |  | P: D: S:
1 | REQ15 | See 1-5 |  |  |  |  | P: D: S: |  |  |  | P: D: S:
1 | REQ17A1-A4 | calculations of values in tables | calculations wrongly coded | coding error | incorrect relesae of PamChips | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | devlopment procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ20 | see 1-5 |  |  |  |  | P: D: S: |  |  |  | P: D: S:
1 | REQ21A | see 1-5 |  |  |  |  | P: D: S: |  |  |  | P: D: S:
1 | REQ21A | table columns | input not correcty configured in BN | user error | calculation fails | no result | P:2 D:1 S:3 | 6 | standard workflow, runtime check on columns in SP | development / procedurs | P:1 D:1 S:3 | 3 |  | REQ42 |  | 3
1 | REQ22 | check on missing values | code development | wrong calculation | incorrect result | wrong diagnosis | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | tables in report | see 1 |  | incorrect result |  | P: D: S: |  |  |  | P: D: S:
1 | REQ026 (A1-A4) | incorrect data in report | code development | wrong calculation | incorrect result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | mix-up of docx reports | use | user / config |  | wrong diagnosis | P:1 D:2 S:5 | 10 | travell card number and date added to report | Proc check on correct travell card (WI) | P:1 D:1 S:5 | 5 |  | REQ27-32 |  | 5
| REQ27-32 | report contents | code development | contents misisng or incorrect | mix-up | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - System testing | check the final rpeorts while testing | P:1 D:1 S:5 | 5
1 | REQ37-38 | see 1-5 |  |  |  |  | P: D: S: |  |  |  | P: D: S:
1 | REQ40-49 | exceptions | code development | exception is missing | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ60-61 | stored data | code devlpoment | incorrect data stored | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | [6 empty] | P: D: S: |  |  |  | P: D: S:
1 | [6 empty] | P: D: S: |  |  |  | P: D: S:
1 | [6 empty] | P: D: S: |  |  |  | P: D: S:
36 | [6 empty] | P: D: S: |  |  |  | P: D: S:
37 | [6 empty] | P: D: S: |  |  |  | P: D: S:
38 | [6 empty] | P: D: S: |  |  |  | P: D: S:
39 | [6 empty] | P: D: S: |  |  |  | P: D: S:
40 | [6 empty] | P: D: S: |  |  |  | P: D: S:
41 | [6 empty] | P: D: S: |  |  |  | P: D: S:
42 | [6 empty] | P: D: S: |  |  |  | P: D: S:
43 | [6 empty] | P: D: S: |  |  |  | P: D: S:
44 | [6 empty] | P: D: S: |  |  |  | P: D: S:
45 | [6 empty] | P: D: S: |  |  |  | P: D: S:
46 | [6 empty] | P: D: S: |  |  |  | P: D: S:
47 | [6 empty] | P: D: S: |  |  |  | P: D: S:
48 | [6 empty] | P: D: S: |  |  |  | P: D: S:
49 | [6 empty] | P: D: S: |  |  |  | P: D: S:
6 | [6 empty] | P: D: S: |  |  |  | P: D: S:

## Explanation Scoring System
Scale of Probability:
Weight |  | Description
5 | Frequent | The failure is extremely likely or certain to occur. The underlying principle or technique is unreliable. | 1
4 | Probable | The failure is likely to occur. The underlying principle or technique is not completely reliable. | 2020-02-18
3 | Occasional | There is a chance for the failure to occur. The underlying principle is advanced and the technique can be controlled. Human error might be of influence.
2 | Remote | It cannot be excluded that the failure occurs.
1 | Improbable | Software Manager | Faris Naji
|  | Head DAS | Rik de Wijn
Scale of Severity: |  |  | software module dasrunqc
Weight |  | Description
5 | Catastrophic | The effect is extremely critical. There is a proven risk to safety (e.g. death or permanent injury), total loss of product or non-compliance to laws.
4 | Critical | The effect is critical. The risk to safety is proven (e.g. life-threatening injury or permanent impairment); total loss of product might be expected, as is non-compliance to laws.
3 | Serious | There is an effect on quality, which can be controlled or repaired or an effect on safety that might require professional medical intervention. There may be some restrictions in handling or the effectiveness of the product or procedure.
2 | Minor | A very modest effect on quality may be expected, which can easily be controlled or repaired, or a minor effect on safety (no professional medical intervention required). The restrictions in handling or lack of effectiveness of the product or procedure are minimal.
1 | Negligible | A negligible or no effect on quality is expected. No or negligible effect on safety (e.g. only inconvenience or temporary discomfort).
Scale of Detectability:
Weight | Description
2 | The failure is impossible to detect or there is no method to routinely test for its occurrence.
1 | Detection of the failure is certain, either visual, or by inspection.
Risk Acceptability Matrix
A) Risks with no or unknown detectability (D=2)
unacceptable risk
undesirable risk
acceptable risk
B) Detectable risks (D=1)
2 | 2 |  | none
3 | 03A |  | programming error | code | incorrect results calculated | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | input/putput testing | unit test | P:1 D:1 S:5 | 5
4 | 5 |  | none |  |  |  |  | 10
5 | 11A and 11B |  | none
6 | 014A and 014B
7 | 017A1 | data | output has incorrect format | code | incorrect results calculated | worst case: wrong diagnosis | P:1 D:2 S:5 | 10 | unit testing | add a unit test | P:1 D:1 S:5 | 5
8 | 018A11, 018A12, 018A13,018A16 |  | programming error | code | incorrect results calculated | wrong diagnosis | P:2 D:2 S:5 | 20 | unit testing | add a unit test | P:1 D:1 S:5 | 5
9 | REQ011A | concordance result | design is not suitable for cioncordance calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
10 | REQ011B | CV result | design is not suitable for CV calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
11 | REQ014A1 | design of simple table | no risk |  |  |  | P: D: S: |  |  |  | 1
12 | REQ014A2 | design of simple trable | no risk | [7 empty] | 5
13 | REQ014A3 | design of simple trable | no risk
14 | REQ014A4 | design of simple trable | no risk
1 | REQ15 | See 1-5
1 | REQ17A1-A4 | calculations of values in tables | calculations wrongly coded | coding error | incorrect relesae of PamChips | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | devlopment procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ20 | see 1-5
1 | REQ21A | see 1-5
1 | REQ21A | table columns | input not correcty configured in BN | user error | calculation fails | no result | P:2 D:1 S:3 | 6 | standard workflow, runtime check on columns in SP | development / procedurs | P:1 D:1 S:3 | 3 |  | REQ42 |  | 3
1 | REQ22 | check on missing values | code development | wrong calculation | incorrect result | wrong diagnosis | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | tables in report | see 1 |  | incorrect result
1 | REQ026 (A1-A4) | incorrect data in report | code development | wrong calculation | incorrect result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | mix-up of docx reports | use | user / config |  | wrong diagnosis | P:1 D:2 S:5 | 10 | travell card number and date added to report | Proc check on correct travell card (WI) | P:1 D:1 S:5 | 5 |  | REQ27-32 |  | 5
| REQ27-32 | report contents | code development | contents misisng or incorrect | mix-up | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - System testing | check the final rpeorts while testing | P:1 D:1 S:5 | 5
1 | REQ37-38 | see 1-5
1 | REQ40-49 | exceptions | code development | exception is missing | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ60-61 | stored data | code devlpoment | incorrect data stored | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1
1
1
6

## Change History
dFMEA Change History
version | reference to row number | changes / additions | 1
2020-02-18
|  | Software Manager | Faris Naji
|  | Head DAS | Rik de Wijn
software module dasrunqc
Rik de Wijn
Faris Naji
2 | 2 |  | none
3 | 03A |  | programming error | code | incorrect results calculated | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | input/putput testing | unit test | P:1 D:1 S:5 | 5
4 | 5 |  | none |  |  |  |  | 10
5 | 11A and 11B |  | none
6 | 014A and 014B
7 | 017A1 | data | output has incorrect format | code | incorrect results calculated | worst case: wrong diagnosis | P:1 D:2 S:5 | 10 | unit testing | add a unit test | P:1 D:1 S:5 | 5
8 | 018A11, 018A12, 018A13,018A16 |  | programming error | code | incorrect results calculated | wrong diagnosis | P:2 D:2 S:5 | 20 | unit testing | add a unit test | P:1 D:1 S:5 | 5
9 | REQ011A | concordance result | design is not suitable for cioncordance calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
10 | REQ011B | CV result | design is not suitable for CV calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
11 | REQ014A1 | design of simple table | no risk |  |  |  | P: D: S: |  |  |  | 1
12 | REQ014A2 | design of simple trable | no risk | [7 empty] | 5
13 | REQ014A3 | design of simple trable | no risk
14 | REQ014A4 | design of simple trable | no risk
1 | REQ15 | See 1-5
1 | REQ17A1-A4 | calculations of values in tables | calculations wrongly coded | coding error | incorrect relesae of PamChips | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | devlopment procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ20 | see 1-5
1 | REQ21A | see 1-5
1 | REQ21A | table columns | input not correcty configured in BN | user error | calculation fails | no result | P:2 D:1 S:3 | 6 | standard workflow, runtime check on columns in SP | development / procedurs | P:1 D:1 S:3 | 3 |  | REQ42 |  | 3
1 | REQ22 | check on missing values | code development | wrong calculation | incorrect result | wrong diagnosis | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | tables in report | see 1 |  | incorrect result
1 | REQ026 (A1-A4) | incorrect data in report | code development | wrong calculation | incorrect result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | mix-up of docx reports | use | user / config |  | wrong diagnosis | P:1 D:2 S:5 | 10 | travell card number and date added to report | Proc check on correct travell card (WI) | P:1 D:1 S:5 | 5 |  | REQ27-32 |  | 5
| REQ27-32 | report contents | code development | contents misisng or incorrect | mix-up | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - System testing | check the final rpeorts while testing | P:1 D:1 S:5 | 5
1 | REQ37-38 | see 1-5
1 | REQ40-49 | exceptions | code development | exception is missing | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ60-61 | stored data | code devlpoment | incorrect data stored | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1
1
1
6

## Approval Page
Signatures and approval
Role | Function title | Name | Signature | Release date
Author | Program Manager IVD | Kristiane Schmidt | 1
Reviewer | Head of DAS | Rik de Wijn | 2020-02-18
Reviewer | Head of BioChip Department | Iskandar Gandasasmita
Authoriser | Managing Director | John Groten
Quality Assurance | QARA | Software Manager | Faris Naji
|  | Head DAS | Rik de Wijn
History |  |  | software module dasrunqc
Version | Date | Remarks/Changes |  | Document Owner
1 | 2018-11-30 | First version of template |  | Theo van der Leij
1.1 | 2019-01-23 | Correction of title and added reference to SOP4500 / POL4500 for detailed definitions |  | Theo van der Leij
1.2 | 2019-06-18 | See CR-19-52. Addition of new sheet for explanation of scoring criteria, use same scales as in POL4500. Clarification of instructions for use of FMEA template. Changed order of columns to match with process. Added new column “Residual Risk & Acceptability”. |  | Theo van der Leij
Faris Naji
2 | 2 |  | none
3 | 03A |  | programming error | code | incorrect results calculated | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | input/putput testing | unit test | P:1 D:1 S:5 | 5
4 | 5 |  | none |  |  |  |  | 10
5 | 11A and 11B |  | none
6 | 014A and 014B
7 | 017A1 | data | output has incorrect format | code | incorrect results calculated | worst case: wrong diagnosis | P:1 D:2 S:5 | 10 | unit testing | add a unit test | P:1 D:1 S:5 | 5
8 | 018A11, 018A12, 018A13,018A16 |  | programming error | code | incorrect results calculated | wrong diagnosis | P:2 D:2 S:5 | 20 | unit testing | add a unit test | P:1 D:1 S:5 | 5
9 | REQ011A | concordance result | design is not suitable for cioncordance calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
10 | REQ011B | CV result | design is not suitable for CV calculation | user error | Incorrect result | worst case: wrong diagnosis | P:2 D:2 S:5 | 20 | runtime - development | standard/reviews designs are used for both routine (WI)and study (study protocol), crucial design parameters should be clear from the report (allowing check) | P:1 D:1 S:5 | 5 | procedure / development | REQ14A3- REQS70A1-A2-A4 - REQ71
11 | REQ014A1 | design of simple table | no risk |  |  |  | P: D: S: |  |  |  | 1
12 | REQ014A2 | design of simple trable | no risk | [7 empty] | 5
13 | REQ014A3 | design of simple trable | no risk
14 | REQ014A4 | design of simple trable | no risk
1 | REQ15 | See 1-5
1 | REQ17A1-A4 | calculations of values in tables | calculations wrongly coded | coding error | incorrect relesae of PamChips | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | devlopment procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ20 | see 1-5
1 | REQ21A | see 1-5
1 | REQ21A | table columns | input not correcty configured in BN | user error | calculation fails | no result | P:2 D:1 S:3 | 6 | standard workflow, runtime check on columns in SP | development / procedurs | P:1 D:1 S:3 | 3 |  | REQ42 |  | 3
1 | REQ22 | check on missing values | code development | wrong calculation | incorrect result | wrong diagnosis | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | tables in report | see 1 |  | incorrect result
1 | REQ026 (A1-A4) | incorrect data in report | code development | wrong calculation | incorrect result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ026 (A1-A4) | mix-up of docx reports | use | user / config |  | wrong diagnosis | P:1 D:2 S:5 | 10 | travell card number and date added to report | Proc check on correct travell card (WI) | P:1 D:1 S:5 | 5 |  | REQ27-32 |  | 5
| REQ27-32 | report contents | code development | contents misisng or incorrect | mix-up | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - System testing | check the final rpeorts while testing | P:1 D:1 S:5 | 5
1 | REQ37-38 | see 1-5
1 | REQ40-49 | exceptions | code development | exception is missing | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1 | REQ60-61 | stored data | code devlpoment | incorrect data stored | wrong result | wrong diagnosis | P:2 D:2 S:5 | 20 | Input - output testing - Unit testing - system testing | development procedures | P:1 D:1 S:5 | 5 |  |  |  | 5
1
1
1
6