## STM
Requirements Traceability Matrix
|  | Document | dascombat-6301 Tracebility Matrix
|  | Version | v1.0
|  | Date | 2020-12-23
|  | Page | 1 of 1
Requirements |  |  |  |  | Risk Management |  |  |  | Test Result
ID# | Description |  | System Requirements | Sofrware Requirements | Risk Item Identification | Hazarad Situation Yes/No | Software Cause | Control Measure | Unit Test Case | Integration Test Case | System Test Case (step # in ?) | Unit test Review | Integration test review | System Test Review
1 | 2 |  | Not Applicable | DAS-COMBAT accepts a matrix that represent the phosphorylation signal |  | no | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply input validation & fit input validation | NA | NA | OK | NA | NA
2 | 003A1 |  | Not Applicable | The results for DAS-COMBAT is a batch corrected version of the input data table (Req 002), | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | test applyModel output size & dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
3 | 5 |  | Not Applicable | Exceptions raised by the DAS-COMBAT will be transmitted to the DAS-COMBAT via the R PACKAGE ENVIRONMENT. | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply input validation & fit input validation | NA | NA | OK | NA | NA
4 | 011A |  | Not Applicable | SP calculates a batch correction model, i.e. fits a model | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
5 | 011B |  | Not Applicable | SP calculates a batch correction i.e. applies a model | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
11 | 014A1 |  | Not Applicable | Table 1 is called “data table” and defined in req21A1 | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx |  | NA | NA | NA | NA | NA
12 | 014A2 |  | Not Applicable | Table 2 is called “aFITmodel data table” and defined in req21A2 | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx |  | NA | NA | NA | NA | NA
14 | 018A11 |  | Not Applicable | The batch model is defined for: No reference, mean.only = FALSE | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
15 | 018A12 |  | Not Applicable | The batch model is defined for: No reference, mean.only = TRUE | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
16 | 018A13 |  | Not Applicable | The batch model is defined for: ref.batch, mean.only = FALSE | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
18 | 018A14 |  | Not Applicable | The batch model is defined for: ref.batch, mean.only = TRUE | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
19 | 018A16 |  | Not Applicable | The batch model is applied to data from batches included in the fit | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply with new batch variable | NA | NA | OK | NA | NA
22 | 22 |  | Not Applicable | Complete arrays only. The DAS-COMBAT checks if there are any missing values | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply input validation & fit input validation | NA | NA | OK | NA | NA
24 | 37 |  | Not Applicable | Any exception returns an informational error message to the DAS-COMBAT via the R PACKAGE ENVIRONMENT. | Class A | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
25 | 38 |  | Not Applicable | Any exception will result in the execution of the DAS-COMBAT being stopped and the error is non-recoverable. | Class A | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply input validation & fit input validation | NA | NA | OK | NA | NA
26 | 40 |  | Not Applicable | An exception is raised when the input matrix contains any missing values. The informational message is “Missing values are not allowed” | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply input validation & fit input validation | NA | NA | OK | NA | NA
28 | 60 |  | Not Applicable | For “fit” with REF (see req18A13, req18A14) the value of REF must occur in batchFactor | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | fit.ref.batch.not.found | NA | NA | OK | NA | NA
29 | 61 |  | Not Applicable | For “apply” with FitTable (see req18A15) the values in batchFactor require to be in batchID in the FitTable data (req21A2) | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | apply with new batch variable | NA | NA | OK | NA | NA
30 | 70 |  | Not Applicable | SP requires passing a test calculation comparing the result to known output. | Class C | No | Yes | dascombat-FRM4503-FMEA sheet.xlsx | dascombat::fit vs sva::ComBat | NA | NA | OK | NA | NA
Note on K10: Silviya Della Chiave: - Risk Avoidance - Risk Transfer - Risk Mitigation - Risk Acceptance